AERIAL CAMERA IN SITU METRIC CALIBRATION GUIDELINES

[Draft, March 10, 2012]

1. Purpose

These guidelines are intended to serve as guidance for the airborne in situ metric calibration of vertically oriented digital and film-based frame type cameras.

The guidelines describe the camera airborne system consisting of the camera, mount, GPS, aircraft, and their spatial relationships.  Calibration results will include both interior and exterior orientation parameter values along with estimates of accuracy.  Results can support validation of geospatial accuracies of imagery products. They guidelines also describe pre-flight, flight and post-flight processes.  Refinements in the guidelines may evolve with future developments.

A benchmark document by Eisenhart (1963), then at the US Bureau of Standards, provides guidance for calibration of measurement systems.  Preparation of guidelines described here was guided by Eisenhart’s concept of measurement system calibration.  When comparisons are made to laboratory methods of calibration, significant improvements in geospatial accuracies are evident.

When geospatial accuracies are specified, they will be in terms defined by federal standards (FGDC-STD-007.2-1988).

2. Conformance

No conformance requirements are established for these guidelines.

3. References

Eisenhart, Churchill (1963) “Realistic Evaluation of the Precision and Accuracy of Instrument Calibration Systems”, Journal of Research of the National Bureau of Standards, Vol. 67C, No. 2, April/June 1963.

FGDC-STD-007.2-1988 (1988) “Geospatial Positioning Accuracy Standards, Part 2:  Standard for Geodetic Networks”, Federal Geographic Data Committee.

Note:  For details regarding theory, history, and extensive reference list, that form the basis of development of these Guidelines for metric calibration of the aerial camera, reference is made to [(Manual of Photogrammetry, Camera Calibration Addendum, 2014)].

4. Authority

The responsible organization for preparing, maintaining, and coordinating work on this guideline is the American Society for Photogrammetry and Remote Sensing (ASPRS), Camera Calibration Committee.  For further information, contact:

Dean C. Merchant, PhD/ merchant.2@osu.edu

5. Terms and Definitions

push-broom – The design of camera imaging devices that gather imagery by means of linear arrays of sensors (the “broom”).  As the transverse array of sensors passes over the image formed by the camera lens, the elements of light energy are recorded, organized and stored.

For additional terms and definitions, reference is made to the “Manual of Photogrammetry, Fourth Edition”, American Society of Photogrammetry, (1980).

6. Symbols, abbreviated terms, and notations

a/c – aircraft

asl – above sea level

agl – above ground level

s/n  - serial number

IFOV – the angle subtended by one pixel

GSD – distance measured at the nadir point subtended by one image pixel

GPS – Global Positioning System

SMAC – Simultaneous Multi-Camera Analytical Calibration

7. Specific Requirements

Airborne cameras considered by this guideline are limited to the frame type, either digital or film-based images.  The digital camera arrays may be treated either as individual arrays or as some form of merging of arrays into a common frame.  The linear array [“push- broom”] configuration is not considered by these guidelines.

Observations of aircraft position relative to the calibration field should be made by a GPS receiver collecting at least L1 and L2 frequencies in a manner useful for post processing.  External event markers, provided by the camera, should accurately time each exposure event and pass the time to the GPS receiver.  Timing accuracy will depend on the ground speed but is typically within 0.0001 seconds.

The aircraft is modified to provide a camera port, either open or with a suitable window.

The GPS antenna should be mounted in the vicinity of the vertical projection of the camera’s optical axis, limited by the structural requirements of the aircraft.  It is preferred, but not required, to use a twin-engine aircraft to avoid possible deformation of the air within the camera’s field of view due to engine cooling air stream and exhaust.

7.1 Pre-Arrival [before the aircraft and camera system and crew arrive at the range airport]

The data provider is expected to instrument the aircraft with camera, GPS, and support equipment.  At this stage, it will be useful for the provider to discuss this issue either with the calibration range coordinator or directly with the camera factory personnel.

In addition to an approved camera system installation, the provider should determine, with the aid of the camera manufacturer, the location of the entrance nodal point with reference to some tangible point on the camera.  Finally, the provider should install a simple temperature measurement device with probe located in the proximity of the entrance node of the lens.

7.2 After Arrival [after aircraft arrives at the range airport]

Measurements should be made to assure the recorded GPS data is accurately related to the crossroad target system of coordinates. A bias may exist between the provider’s GPS measurements and the cross road range coordinate system. If a bias exists, it will be applied during data processing as a correction to the GPS-determined geospatial coordinates of each exposure station.

Finally, a preflight briefing of the crew should be held informing them of the flight pattern over the range and, to a limited extent, the theory of the in situ procedure.

7.3 Flight Mission

Image and data collection over the calibration range should be conducted as closely as possible to conditions and procedures expected during a typical operational mission.  It is understood that some variations, for instance in altitude, may be accepted (Eisenhart, 1963).   During the photo mission, GPS observations must be collected at a base station located within a specified distance of the range center. For example, a maximum of 15 kilometers (9.3 miles) distance, base to range, is specified for a Camera Class II Calibration [see Annex A].

During collection of calibration imagery, the nominal pitch angle of the aircraft should be measured and recorded for subsequent applications. It is the angle that will be used for aircraft stabilization during ground measurements of phase center to entrance node offsets.  It is also the pitch angle to be used during all subsequent flight operations, particularly when large scale imagery is collected.

For a Camera Class II calibration [see Annex A], the flight pattern will consist of a minimum of six lines, each acquiring a single frame over the intersection of roads.  Given (not required) a range where the crossroads are oriented to north and east, the nominal flight azimuths will be one each at:

north / south / 30 / 45 / 60 / 90 degrees

These directions are arbitrarily chosen as the minimum number to balance the coverage of target images across the image field and to suppress possible remaining systematic errors not accounted for in the mathematical model.  For cameras with a non-square format, some adjustments to the azimuths may be necessary to cause images to appear in the extreme corners of the image format.

7.4 Post Mission [before the provider’s crew leaves the range airport]

Spatial offsets between the phase center of the antenna and entrance node of the camera are to be measured.  It is essential that these offsets be measured in a coordinate system parallel to the camera coordinate system and to accuracy within one centimeter along each axis.  For these measurements, the aircraft must be stabilized with the wings level and the pitch angle established equal to that measured during the calibration flight.  The camera should be set to level and the remaining mount angles set to zero. These settings facilitate offset measurements along camera axes.

After the flight, it should be assured that the mission has collected adequate and reliable imagery and GPS data.  These procedures, although they may be time consuming, will assure that data processing may proceed after that and no return of the aircraft/crew will be needed.

Comments from the provider’s crew will be useful in refining all field aspects of the calibration process.

7.5 Data Processing and Reporting

Measurement of imagery and processing of the GPS data and range coordinates may be conducted by commercially available or other software.

For the final step, a report of calibration must be produced and distributed for subsequent applications.  The report should include the specifications describing all equipment and procedures when pertinent (camera s/n and type, optical filter, film, magazine s/n, processing, scanning, flight height agl, asl, temperature at the lens and a/c cabin, and antenna to nodal offsets).

The mathematical model used to represent the camera’s interior orientation will depend on the level of correction applied to the image provided to the user.  In the case of the film-based camera, only the scanned image without corrections is normally provided to the user.  In this case, the conventional SMAC model is to be evaluated resulting in estimates of interior orientation including focal length, principle point coordinates and parameters describing radial and decentering distortions.

In the event the digital camera is used, generally, the influences of all but the focal length and principal point coordinates need to be treated as unknown parameters for adjustment purposes. In concept, all those parameters not corrected for in the image provided to the user should be treated as unknown parameters in the calibration computations.

In all cases, the calibration report must state the mathematical model used to represent the interior orientation of the camera, the estimates of parameter values of the model and estimates of errors in the adjusted parameters.

Finally, the source and characteristics of thermal corrections should be indicated when available.

For film-based cameras, the average for all frames of the rmse fit of the observations of fiducial marks to their published values should be provided in the report.  This will assure the reliability of fiducial mark stability and data.

Typically, calibration adjustment computation results demonstrate a root mean square error (rmse) of image residuals equal to or less than one-half pixel and should be indicated in the final report.

Annex A

Crossroad Camera Calibration Range Specifications

The range geospatial geometry is patterned after the USGS bank collimator method used for routine laboratory calibration.  However, the airborne approach allows the camera system geometry to be measured under conditions of environment that exist during flight operations.  The notion of calibration using operational conditions is termed “in situ calibration”.  This method has been shown to be substantially more accurate than a laboratory approach since all environmental factors are involved in the calibration process. The in situ approach to calibration can be done with greater convenience for the producers and at reduced cost.

The calibration range is built on any cross-roads provided the intersection is a nominal 90 degrees.  Roads must have visibility from the air for a distance from the intersection at least equal to the flight height above ground intended for the calibration flights.  An additional 5 % extension of the range, with targets, should be included to allow for assured target image coverage.

Target distribution is spaced nominally at 7.5 degree intervals along the roads.  With the exceptions of the outer several targets of each leg, the spacing variations between targets may be as great as 50 percent.  This will permit a good site for GPS ground survey purposes.

The following design specifications are based on the assumptions of image pixel size of 14 microns (typical for film-based imagery), a camera with 152.4 mm (6 inch) focal length and a flight height of approximately 600 meters (2000 feet) above ground level (AGL).

- White circular targets of 305 mm (1.0 foot) diameter centered on a black target of 366 mm (1.2 foot) diameter (A white target on a black road surface will approximate this requirement.)

- Nominal spacing of targets at 62.8 meters (260 feet) (Note 50% variation of spacing above.)

- Add at least two targets at +/- 31.4 meters (130 feet) on either side of the last outboard target on each leg.

- Control survey should result in “XYZEC”  [X, Y, Z, Earth Centered] coordinates in a system and geodetic datum to be used to position the aircraft [WGS84]

- Relative spatial accuracy along each axis should be within 20 mm (0.8 inches) standard error with respect to the base station.

For flights of higher or lower altitudes, the linear dimensions should be increased or decreased accordingly.  Considerations of focal length, pixel size, and camera field angle as they influence target size and distribution should also be made.